General ideas for next version(s)

- White and gray ferrules don’t fit on wire; they are for a thinner wire insulation thickness. Maybe get thinner wire, I think multi-color wire and having one color ferrule we use with all the wire colors is the way to go.

- Get resistors on the PCB and a rotary switch.

- Make the box a bit shorter,

- Put the power plug on the side up against the edge of the box.

- Put a status LED on the front that is maybe tied to one of the square waves.

- Add a rotary dial for the bi-phasic speed.

- Remove the version number from the 3D printed boxes.

- Add a 2-probe/4-probe switch to the PCB.

- Looks like AI0-3 connect to the CDAQ ports 7-4 (in that order), we need to fix these.

- Move components around so a screwdriver can get to the D-sub screws.

- Add an SN line to the silk screen, so we can add SNs in paint pens.

- Drop next version to 5V, bypass the Arduino linear regulator but add some caps by the 5V line to help isolate it from the switching noise.